# DNA sequencing results: Noisy data

Background noise appears as many smaller, undefined peaks under your sequence peaks of interest. This noise is always present, but with well-prepared samples of good signal strength, it will be undetectable.

If the noises are too strong, the software has difficulty in resolving the base peaks from background noise. This results in poor sequence reliability

Example of noisy data: presence of Background noise misleading the automated base caller at various points in the chromatogram

Causes of Noisy Sequence

- Dirty DNA

- Low DNA concentration. As the signal intensity is low, the background noise will be artificially amplified by the software in order to called bases

- Inadequate primers:

- Inefficient primer annealing

- Multiple priming sites

- Noisy data caused by random priming of a N-1 primer synthesis

Possible solutions

- Increase template concentration

- Re-do DNA prep

- Request alternative sequencing programmes

- Choose another primer